## Slide 1
Data Analysis
Philips monitoring alarm data

## Slide 2
High Priority Arrhythmia (***)
Asystole
Medium Priority Arrhythmia (**)
Non-sustained Vtach
High Priority Parameter (!!!)
Apnea
Inop
(some can be configured a higher priority)
Medium Priority Parameter (!!)
Vfib/Tach
Extreme Tachycardia
Extreme Bradycardia
V-Tach
SVT
SpO2 Desat
Invasive Pressure Line Disconnect
Invasive Pressure Extreme Limit
High or Low Limit Violations:
SpO2
Respiratory
NBP
Invasive Pressure
Temp
CO2
Other
Leads Off
Replace Battery
Cannot analyze ECG
Others outlined in IFU; not captured by PIIC iX
Ventricular Rhythm
Bigeminy
Run PVCs
Pair PVCs
Trigeminy
PVCs per minute
Multiform PVCs
Pause
Pacer not capture
Pacer not pace
Missed beat
HR High/Low
Atrial Fib/Atrial Fib End
Irregular Heart Rate/Irregular Heart Rate End
Philips Monitoring Alarm Categories
DONE

## Slide 3
Alarm Totals by Category Units
Medium priority arrhythmia alarms contribute to more than half of all the alarms – despite many being defaulted off
Data Analytics
Chart: 
| Red Arrhythmia | Yellow Arrhythmia | Red Parameter | Yellow Parameter | Inop/Technical
Totals | 34607 | 304009 | 33115 | 318442 | 74930
Total: 765,103
Red Text
Notes: This is due to the HR high and low limits

## Slide 4
Total Monitoring Alarms Per Unit
Overall disruption and noise on the units
Total: 765,103
Chart: 
| ART CVICU | Main Neuro | PCICU | Main Tele 9-Neuro | MUSC ED | ART Tele 3-Cardio Inpatient | ART Tele A5E | ART Tele A5W | ART MSICU | ART CPC | STICU | MICU | ART Tele 4-CTSurgery | ART Tele 7-Oncology | Main 6 West | ART HVC PP | Main Tele 8-Medsurg | ART Tele 6-Digestive | Main Tele 10-Ortho | ART PACU | Main Tele 6-Renal | Main Tele 7-West | Main Tele 6-Trauma | Adult IR | ART Holding | Main 7 W | Main Radhold | Main Tele 5-Obs | Main LD | Main 6 East | ART Tele 3-CathPrep | ART Tele 1-Holding | Dialysis | ART Tele Telemetry | Main Tele 2-EAU | ART Tele 3-MSICU
Total | 83581 | 71525 | 69836 | 67694 | 61478 | 49772 | 42114 | 40991 | 40798 | 35860 | 32177 | 27089 | 26324 | 17547 | 17188 | 12609 | 12371 | 11532 | 9302 | 8394 | 7132 | 4703 | 3873 | 3075 | 2561 | 1800 | 1161 | 1125 | 636 | 420 | 179 | 107 | 96 | 34 | 14 | 5
Data Analytics
DONE

## Slide 5
Total Alarms Per Patient Bed Per Day
Benchmark to compare and evaluate change impact.
Unit 54 has the highest overall.
It is one of the units that shifts/adds on to an MTs load for part of the day.
Data Analytics
Chart: 
| ART CVICU | Main Neuro | PCICU | Main Tele 9-Neuro | MUSC ED | ART Tele 3-Cardio Inpatient | ART Tele A5E | ART Tele A5W | ART MSICU | ART CPC | STICU | MICU | ART Tele 4-CTSurgery | ART Tele 7-Oncology | Main 6 West | ART HVC PP | Main Tele 8-Medsurg | ART Tele 6-Digestive | Main Tele 10-Ortho | ART PACU | Main Tele 6-Renal | Main Tele 7-West | Main Tele 6-Trauma | Adult IR | ART Holding | Main 7 W | Main Radhold | Main Tele 5-Obs | Main LD | Main 6 East | ART Tele 3-CathPrep | ART Tele 1-Holding | Dialysis | ART Tele Telemetry | Main Tele 2-EAU | ART Tele 3-MSICU
Total | 139.9 | 149.5 | 172.7 | 94.8 | 85.9 | 88.3 | 139.2 | 83.5 | 79.2 | 103.6 | 72.4 | 54.6 | 61.9 | 67.1 | 43.2 | 21.3 | 42.7 | 50.2 | 67.6 | 30.8 | 68 | 65.5 | 42.6 | 14.1 | 12.5 | 20.1 | 7.2 | 43.7 | 13.4 | 7.1 | 44.8 | 53.5 | 4.2 | 3.2 | 7 | 5
Red Text

## Slide 6
Total Alarms Per Patient Bed Per Day
Comparison against published metrics.
Best practice of having many medium priority arrhythmia alarms defaulted off has had an impact
Published study results
Advocate Units
Chart: 
| Drew Study | Johns Hopkins A | Johns Hopkins B | ART CVICU | Main Neuro | PCICU | Main Tele 9-Neuro | MUSC ED | ART Tele 3-Cardio Inpatient | ART Tele A5E | ART Tele A5W | ART MSICU | ART CPC | STICU | MICU | ART Tele 4-CTSurgery | ART Tele 7-Oncology | Main 6 West | ART HVC PP | Main Tele 8-Medsurg | ART Tele 6-Digestive | Main Tele 10-Ortho | ART PACU | Main Tele 6-Renal | Main Tele 7-West | Main Tele 6-Trauma | Adult IR | ART Holding | Main 7 W | Main Radhold | Main Tele 5-Obs | Main LD | Main 6 East | ART Tele 3-CathPrep | ART Tele 1-Holding | Dialysis | ART Tele Telemetry | Main Tele 2-EAU | ART Tele 3-MSICU
Total | 187 | 215 | 317 | 140 | 150 | 173 | 95 | 86 | 88 | 139 | 84 | 79 | 104 | 72 | 55 | 62 | 67 | 43 | 21 | 43 | 50 | 68 | 31 | 68 | 66 | 43 | 14 | 13 | 20 | 7 | 44 | 13 | 7 | 45 | 54 | 4 | 3 | 7 | 5
Data Analytics
Red Text

## Slide 7
Chart: 
| ART Tele | MAIN Tele | ART CVICU | Main Neuro | PCICU | MUSC ED | ART MSICU | ART CPC | STICU | MICU | Main 6 West | ART HVC PP | ART PACU | Adult IR | ART Holding | Main 7 W | Main Radhold | Main LD | Main 6 East | Dialysis
Total | 188605 | 106214 | 83581 | 71525 | 69836 | 61478 | 40798 | 35860 | 32177 | 27089 | 17188 | 12609 | 8394 | 3075 | 2561 | 1800 | 1161 | 636 | 420 | 96
Total Monitoring Alarms Per Unit With CTC
MT alarm load based on current unit assignments – one MT compared to multiple RNs addressing alarms
Data Analytics
Total: 765,103
Alarm totals per CTC stations
Red Text

## Slide 8
Chart: 
| ART Tele 3-Cardio Inpatient | ART Tele A5E | ART Tele A5W | ART Tele 4-CTSurgery | ART Tele 7-Oncology | ART Tele 6-Digestive | ART Tele 3-CathPrep | ART Tele 1-Holding | ART Tele Telemetry | ART Tele 3-MSICU | Main Tele 9-Neuro | Main Tele 8-Medsurg | Main Tele 10-Ortho | Main Tele 6-Renal | Main Tele 7-West | Main Tele 6-Trauma | Main Tele 5-Obs | Main Tele 2-EAU
Total | 49772 | 42114 | 40991 | 26324 | 17547 | 11532 | 179 | 107 | 34 | 5 | 67694 | 12371 | 9302 | 7132 | 4703 | 3873 | 1125 | 14
CTC Unit Breakout
MT alarm load based on current unit assignments – one MT compared to multiple RNs addressing alarms
Data Analytics
Alarm totals per CTC stations
Red Text

## Slide 9
Findings
ECG Leads Off is ranked in the top 3 most frequent alarms in 12 of the units
HR High, HR Low, and SpO2 Low are commonly found in the top 3 for many units
Considerations
Addressing ECG Leads Off, HR High, HR Low and SpO2 Low should have an impact across multiple units.
Are the PVC trigger rates in CCU and CV Heart at actionable levels?
ART CPC |  | ART CVICU |  | ART HVC PP |  | ART MSICU |  | ART HOLDING |  | ART PACU | 
SpO2 Low | 7,285 | SpO2 Low | 18,983 | SpO2 Low | 2,750 | SpO2 Low | 9,690 | ECG Leads Off | 1,200 | ECG Leads Off | 2,537
RR Low | 5,774 | ABPm Low | 6,438 | ECG Leads Off | 2,668 | ABPm Low | 3,408 | Pair PVCs | 224 | SpO2 Low | 1,523
ECG Leads Off | 3,881 | Pair PVCs | 5,847 | RR High | 1,455 | ECG Leads Off | 2,209 | SpO2 Low | 175 | RR Low | 1,469
ART TELE 1-HOLDING |  | ART TELE 3-CARD-INPT |  | ART TELE 4-CTSURG |  | ART TELE A5E |  | ART TELE A5W |  | ART TELE 6-DIGESTIVE | 
Run PVCs High | 22 | PVC Rate | 8,148 | PVC Rate | 5,086 | RR High | 13,222 | PVC Rate | 7,511 | ECG Leads Off | 2,746
VTach | 19 | Non-Sust. VT | 5,948 | Non-Sust. VT | 4,201 | SpO2 Low | 6,751 | HR High | 6,226 | HR High | 2,516
Vent Fib/Tach | 15 | HR High | 5,194 | HR High | 3,060 | ECG Leads Off | 4,366 | Non-Sust. VT | 4,562 | Missed Beat | 1,099
ART TELE 7-ONCOLOGY |  | ART TELEMETRY |  | MAIN ADULT IR |  | MAIN 6 EAST |  | MAIN 6 WEST |  | MAIN 7 WEST | 
HR High | 4,678 | PVC Rate | 14 | SpO2 Low | 687 | SpO2 Low | 109 | SpO2 Low | 11,182 | SpO2 Low | 791
ECG Leads Off | 2,204 | ECG Leads Off | 11 | ECG Leads Off | 655 | RR Low | 81 | Pulse High | 2,950 | Pulse Low | 489
PVC Rate | 1,890 | HR Low | 6 | RR Low | 428 | ECG Leads Off | 60 | Desat | 1,167 | Pulse High | 218
MAIN MICU |  | MAIN STICU |  | MAIN LD |  | MAIN NEURO |  | MAIN RADHOLD |  | MAIN TELE 2-EAU | 
SpO2 Low | 5636 | RR Low | 4,873 | SpO2 Low | 180 | ABPs High | 8,947 | ECG Leads Off | 699 | Leads Unplug. | 12
RR High | 2445 | SpO2 Low | 4,499 | Pulse High | 167 | RR High | 7,868 | SpO2 Low | 158 | ECG Leads Off | 1
ABPm Low | 2269 | ABPs High | 3,169 | ABPs High | 154 | SpO2 Low | 6,349 | NBPs Low | 65 | Pause | 1
MAIN TELE 3-CATHPREP |  | MAIN TELE 5-OBS |  | MAIN TELE 6-RENAL |  | MAIN TELE 6-TRAUMA |  | MAIN TELE 7-WEST |  | MAIN TELE 8-MEDSURG | 
HR High | 84 | HR High | 632 | Non-Sust. VT | 1,427 | PVC Rate | 947 | PVC Rate | 877 | ECG Leads Off | 3,311
Extreme Tachy | 40 | ECG Leads Off | 153 | HR High | 1,142 | HR High | 548 | HR High | 737 | HR High | 1,807
ECG Leads Off | 27 | Missed Beat | 92 | PVC Rate | 1,037 | ECG Leads Off | 360 | ECG Leads Off | 607 | Missed Beat | 1,735
MAIN TELE 9-NEURO |  | MAIN TELE 10-ORTHO |  | DIALYSIS |  | MAIN TELE 3-MSICU |  | ED |  | PCICU | 
SpO2 Low | 8,466 | HR High | 2,223 | SpO2 Low | 59 | ECG Leads Off | 2 | SpO2 Low | 12,870 | SpO2 Low | 18,959
ECG Leads Off | 7,031 | Non-Sust. VT | 1,238 | ECG Leads Off | 17 | Replace Batt | 2 | RR High | 9,188 | Desat | 5,464
HR Low | 6,755 | ECG Leads Off | 1,223 | Desat | 11 | Tele Batt Low | 1 | ECG Leads Off | 8,965 | HR High | 5,264
Top 3 Alarms per Unit
Red Text

## Slide 10
Medium Priority Parameter Alarms by > 1000 Frequency
These account for 95% of the medium priority parameter alarms
Data Analytics
Total Number of Medium Priority Parameter Alarms that Occurred Across All Units – Over 30 Days – 318,442
Chart: 
| SpO2 Low | RR High | RR Low | ABPs High | ABPm Low | ABPs Low | NBPs High | ABPm High | Pulse High | awRR High | NBPs Low | etCO2 Low | CVPm High | Pulse Low | ARTm Low | RAPm High | CPP Low | Temp Low | CVPm Low | PAPd High | ARTm High | NBPm High | Unknown Alarm | ARTs Low
Series 1 | 117113 | 52400 | 38450 | 18562 | 13025 | 10137 | 7602 | 6486 | 4742 | 4611 | 4401 | 4261 | 3463 | 3061 | 2884 | 2485 | 1488 | 1327 | 1239 | 1218 | 1175 | 1105 | 1105 | 1085
SpO2 Low accounts for 37% of the total of all the medium priority parameter alarms
DONE

## Slide 11
Chart: Alarm Violation Value
| 87 | 86 | 85 | 84 | 83 | 82 | 81 | 80 | 79 | 78 | 77 | 76 | 75 | 74 | 73 | 72 | 71 | 70 | 69 | 68 | 67 | 66 | 65 | 64-54 | 53-21
Series 1 | 574 | 647 | 554 | 374 | 329 | 245 | 199 | 193 | 153 | 134 | 105 | 71 | 48 | 56 | 35 | 39 | 35 | 19 | 28 | 23 | 18 | 15 | 12 | 55 | 30
Medium Priority SpO2 Parameter Alarms
Literature recommends a default value of 88 for low limit
Total SpO2 violations that occurred in MICU with a limit set to 88 - Over a 30 days period – 3,991
DONE

## Slide 12
Chart: Alarm Count
| ART CVICU | Main Neuro | PCICU | MUSC ED | ART Tele A5E | Main Tele 9-Neuro | STICU | ART MSICU | MICU | ART CPC | Main 6 West | ART HVC PP | ART PACU | Adult IR | Main 7 W | Main LD | ART Holding | Main Radhold | Main 6 East | Dialysis | ART Tele A5W | ART Tele 3-Cardio Inpatient | Main Tele 8-Medsurg | ART Tele 6-Digestive | ART Tele 4-CTSurgery | ART Tele 7-Oncology | Main Tele 10-Ortho | Main Tele 6-Renal | Main Tele 6-Trauma | Main Tele 5-Obs | ART Tele 3-CathPrep | Main Tele 7-West | Main Tele 2-EAU | ART Tele Telemetry | ART Tele 1-Holding | ART Tele 3-MSICU
Alarm Count | 44688 | 44099 | 43658 | 33690 | 25048 | 20937 | 20832 | 19668 | 18036 | 17070 | 15156 | 6626 | 3920 | 1679 | 1502 | 614 | 517 | 386 | 237 | 67 | 12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unique Alarms | 56 | 43 | 44 | 32 | 12 | 17 | 46 | 57 | 50 | 22 | 9 | 13 | 10 | 14 | 4 | 8 | 7 | 15 | 5 | 3 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Medium Priority Parameter Alarm Totals By Unit
Are all parameter alarms in the MSICU and CVICU actionable?
Data Analytics
Total Number of Unique Medium Priority Parameter Alarms By Unit
Over 30 Days – 318,442
Total Unique Parameters Alarms Types by Unit
Total Parameter Alarms by Unit
DONE

## Slide 13
Chart: 
| HR High | PVC Rate | HR Low | Non-Sustain VT | Pair PVCs | Missed Beat | Run PVCs High | Multiform PVCs | Afib | Pacer Not Pacing | Pause | End Afib | Vent Bigeminy | Vent Rhythm | Irregular HR | Pacer Not Capture | R-on-T PVCs | End Irregular HR | Vent Trigeminy
Series 1 | 58060 | 42540 | 32995 | 31601 | 25206 | 18556 | 17895 | 17856 | 12572 | 10064 | 8441 | 6767 | 4912 | 3760 | 3400 | 2541 | 2470 | 2025 | 1472
Medium Priority Arrhythmia Alarms by > 1000 Frequency
These account for 99% of the medium priority arrhythmia alarms
HR High and HR Low accounts for 30% of these alarms: can default limits be safely adjusted?
Data Analytics
Total Number of Medium Priority Arrhythmia Alarms that
Occurred Across All Units – Over 30 Days – 304,009
HR High/Low account for 30% of the total of the medium priority arrhythmia alarms
Red Text

## Slide 14
Data Analytics
High Priority/Red Arrhythmia Alarms
High Priority/Red Parameter Alarms
Total Number of High Priority Monitoring Alarms Across Units – Over 30 Days - 67,722
High Priority Monitoring Alarms by > 100 Frequency
These account for 99% of the high priority monitoring alarms
XXX and XXX account for 49% of the total high priority alarms.
Is an Extreme Brady limit of 50 actionable?
Chart: 
| Desat | VTach | Extreme Tachy | Apnea | Asystole | Extreme Brady | High Paw | Vent Fib/Tach | ABPm High | ABPm Low | ABP Disconnect | ABPs High | Tachy Pulse | ABPs Low | Brady Pulse | Vent Alarm (red parameter) | ICPm High | ART Disconnect
Series 1 | 16817 | 12290 | 9300 | 6166 | 5722 | 5633 | 1850 | 1662 | 1525 | 1383 | 979 | 817 | 806 | 756 | 684 | 607 | 265 | 183
Red Text

## Slide 15
Data Analytics
Comparison of Leads Off to High Priority Alarms
High Priority Alarms
Variation in severity of alarm priority for Leads Off
ECRI recommends High Priority
High Priority/Red Arrhythmia Alarms
High Priority/Red Parameter Alarms
Inop Alarm for ECG Leads off – variation in priority
Chart: 
| ECG Leads Off | Desat | VTach | Extreme Tachy | Apnea | Asystole | Extreme Brady | High Paw | Vent Fib/Tach | ABPm High | ABPm Low | ABP Disconnect | ABPs High | Tachy Pulse | ABPs Low | Brady Pulse | Vent Alarm (red parameter) | ICPm High | ART Disconnect
Series 1 | 65727 | 16817 | 12290 | 9300 | 6166 | 5722 | 5633 | 1850 | 1662 | 1525 | 1383 | 979 | 817 | 806 | 756 | 684 | 607 | 265 | 183
DONE

## Slide 16
Data Analytics
Select Inop Monitoring Alarms
Leads Off is the most common and there is variation in Leads Off priority and priority for a single lead off
Chart: 
| ECG Leads Off - Medium | ECG Leads Off - Low | Check Patient ID - Medium | Replace Telemetry Battery - Medium | Telemetry Battery Low - Low | RA Leads Off - Medium | LL Leads Off - Medium | Leadset Unplugged - Low | Battery Empty - Low | Unknown INOP | Battery Empty - Low - Low | LA Leads Off - Medium | Battery Empty - Medium | Transmitter Off - Low | Check Equipment - Medium | Check ECG Source - Medium | NBP Cuff Overpress - Low | Charge Battery - Low | Replace Tele Battery - Low | PW:Action Required - Medium | Battery Malfunction | Cuff Not Deflated - Low | Insert Battery - Medium
Series 1 | 33995 | 31732 | 3607 | 1927 | 1112 | 669 | 551 | 449 | 410 | 149 | 89 | 65 | 48 | 28 | 27 | 20 | 20 | 19 | 7 | 3 | 1 | 1 | 1
Total Number of InOp Monitoring Alarms That
Occurred Across All Units – Over 30 Days – 74,930
DONE

## Slide 17
Data Analytics
Inop/Technical Alarms by Unit
Chart: 
| MUSC ED | ART CPC | Main Tele 9-Neuro | ART Tele A5E | ART Tele 3-Cardio Inpatient | ART Tele A5W | Main Tele 8-Medsurg | ART CVICU | ART Tele 6-Digestive | ART Tele 4-CTSurgery | ART HVC PP | ART PACU | ART Tele 7-Oncology | PCICU | ART MSICU | Main Neuro | STICU | MICU | Main Tele 10-Ortho | ART Holding | Main Tele 6-Renal | Main Radhold | Main Tele 7-West | Adult IR | Main Tele 6-Trauma | Main Tele 5-Obs | Main 6 East | ART Tele 3-CathPrep | Main 6 West | Dialysis | Main Tele 2-EAU | ART Tele Telemetry | Main 7 W | ART Tele 1-Holding | Main LD | ART Tele 3-MSICU
Series 1 | 9542 | 7596 | 7286 | 4610 | 4266 | 4008 | 3637 | 3140 | 2987 | 2938 | 2681 | 2654 | 2515 | 2406 | 2395 | 2289 | 1775 | 1664 | 1365 | 1224 | 990 | 725 | 699 | 691 | 478 | 191 | 60 | 27 | 21 | 18 | 13 | 11 | 10 | 7 | 6 | 5
Total: 74,930
DONE
Notes: Leads off is the most common for the ED and Units 30 & 40

## Slide 18
Data Analytics
AAMI task force evaluating if 20% of patients cause the majority of the alarms (65%-80%)
Patient Outliers
Can impact total alarms on a unit
Customization may have impact here
Chart: 
| MICU1 | MICU2 | MICU3 | MICU4 | MICU5 | MICU6 | MICU7 | MICU8 | MICU9 | MICU10 | MICU11 | MICU12 | MICU13 | MICU14 | MICU15 | MICU16 | MICU17 | MICU18
Series 1 | 171 | 86 | 145 | 54 | 189 | 31 | 28 | 11 | 28 | 15 | 31 | 42 | 34 | 15 | 31 | 109 | 37 | 60
Room Number
MICU, 7-Mar-2018
DONE

## Slide 19
Data Analytics
Summary of Focus Unit Comparisons
Tele Units top three alarms are the same; some variation in CCU and ED.
A5W
NNICU
MSICU
MICU
40,991
192,411
27,089
40,798
84
210
55
79
Total Alarms
Total Alarms Per Patient Bed Per Day
Top 3 Alarms Overall
1. PVC Rate | 7,511
2. HR High | 6,226
3. Non-Sust. VT | 4,562
1. SpO2 High | 67,924
2. SpO2 Low | 64,480
3. Desat | 26,071
1. SpO2 Low | 5,636
2. RR High | 2,445
3. ABPm Low | 2,269
1. SpO2 Low | 9,690
2. ABPm Low | 3,408
3. ECG Leads Off | 2,209
DONE

## Slide 20
Data Analytics
A5W
Leads Off can be caused by leads becoming unattached unintentionally; or when staff removes because patient goes for a test or is discharged from the ED but the monitor is left on; or staff is monitoring only SpO2 & NBP but leaves ECG on
Chart: 
| PVC Rate | HR High | Non-Sustain VT | ECG Leads Off | Run PVCs High
Series 1 | 7511 | 6226 | 4562 | 3387 | 2866
A5W Top 5 Alarms – Over 30 Days – 24,552
These 5 alarms account for 60% of the total A5W monitoring alarms.
Red Text

## Slide 21
Data Analytics
NNICU
Consider lowering SpO2 low limit to 88 and increasing delay to 15 or 20 seconds – at 90 now with delay of 10 seconds
Chart: 
| SpO2 High | SpO2 Low | Desat | HR Low | ECG Leads Off
Series 1 | 67924 | 64480 | 26071 | 11690 | 5098
NNICU Top 5 Alarms – Over 30 Days – 175,263
These 5 alarms account for 91% of the total NNICU monitoring alarms.
Red Text

## Slide 22
Data Analytics
MICU
Consider whether the PVC Rate limit can be increased from 10 per minute to a higher rate as it is at the other hospitals
Chart: 
| SpO2 Low | RR High | ABPm Low | HR High | RR Low
Series 1 | 5636 | 2445 | 2269 | 1932 | 1867
MICU Top 5 Alarms – Over 30 Days – 14,149
These 5 alarms account for 52% of the total MICU monitoring alarms.
Red Text
Notes: PVC Rate 20 at Good Shep & 30 at South Sub

## Slide 23
Data Analytics
MSICU
Determine if their transport process is contributing to Leads Off alarms
Chart: 
| SpO2 Low | ABPm Low | ECG Leads Off | HR High | Desat
Series 1 | 9690 | 3408 | 2209 | 2209 | 1642
MSICU Top 5 Alarms – Over 30 Days – 19,158
These 5 alarms account for 47% of the total MSICU monitoring alarms.
Red Text

## Slide 24
Children's Data Analysis
Philips monitoring alarm data

## Slide 25
Alarm Totals by Category Units
Medium priority arrhythmia alarms contribute to more than half of all the alarms – despite many being defaulted off
Data Analytics
Chart: 
| Red Arrhythmia | Yellow Arrhythmia | Red Parameter | Yellow Parameter | Inop/Technical
Totals | 15504 | 66277 | 53327 | 257734 | 24590
Total: 417,432
Red Text
Notes: This is due to the HR high and low limits

## Slide 26
Total Monitoring Alarms Per Unit
Overall disruption and noise on the units
Total: 417,432
Chart: 
| CH NNICU | CH SC5 EC | CH 8D | CH PICU | CH 7C | CH SC8 | CH Peds ED | CH PPR | CH Peds Same D
Total | 192411 | 72803 | 45204 | 30371 | 26018 | 23287 | 22747 | 2573 | 2018
Data Analytics
DONE

## Slide 27
Total Alarms Per Patient Bed Per Day
Benchmark to compare and evaluate change impact.
Unit 54 has the highest overall.
It is one of the units that shifts/adds on to an MTs load for part of the day.
Data Analytics
Chart: 
| CH NNICU | CH SC5 EC | CH 8D | CH PICU | CH 7C | CH SC8 | CH Peds ED | CH PPR | CH Peds Same D
Total | 209.8 | 88.7 | 92 | 104.7 | 93.5 | 86.6 | 69.1 | 14.3 | 33.2
Red Text

## Slide 28
Total Alarms Per Patient Bed Per Day
Comparison against published metrics.
Best practice of having many medium priority arrhythmia alarms defaulted off has had an impact
Published study results
Advocate Units
Chart: 
| Drew Study | Johns Hopkins A | Johns Hopkins B | CH NNICU | CH SC5 EC | CH 8D | CH PICU | CH 7C | CH SC8 | CH Peds ED | CH PPR | CH Peds Same D
Total | 187 | 215 | 317 | 210 | 89 | 92 | 105 | 94 | 87 | 69 | 14 | 33
Data Analytics
Red Text

## Slide 29
Findings
ECG Leads Off is ranked in the top 3 most frequent alarms in 12 of the units
HR High, HR Low, and SpO2 Low are commonly found in the top 3 for many units
Considerations
Addressing ECG Leads Off, HR High, HR Low and SpO2 Low should have an impact across multiple units.
Are the PVC trigger rates in CCU and CV Heart at actionable levels?
7C |  | 8D |  | NNICU |  | PEDS ED |  | PEDS SAME DAY | 
SpO2 Low | 8,888 | SpO2 Low | 13,997 | SpO2 High | 67,924 | RR High | 8,347 | RR High | 418
ECG Leads Off | 3,800 | Desat | 7,465 | SpO2 Low | 64,480 | SpO2 Low | 2,694 | SpO2 Low | 332
RR High | 3,006 | HR High | 7,263 | Desat | 26,071 | ECG Leads Off | 2,567 | etCO2 Low | 208
Top 3 Alarms per Unit
PICU |  | PPR |  | SC5 EC |  | SC8 | 
SpO2 Low | 6,060 | ECG Leads Off | 725 | SpO2 Low | 37,070 | SpO2 Low | 12,488
RR High | 3,652 | RR High | 377 | HR Low | 12,733 | Desat | 3,101
HR Low | 2,895 | SpO2 Low | 322 | Desat | 10,211 | SpO2 High | 1,674
Red Text

## Slide 30
Medium Priority Parameter Alarms by > 100 Frequency
These account for 99% of the medium priority parameter alarms
Data Analytics
Total Number of Medium Priority Parameter Alarms that Occurred Across All Units – Over 30 Days –257,734
Chart: 
| SpO2 Low | SpO2 High | RR High | RR Low | ABPm High | NBPs High | ABPm Low | ARTm High | etCO2 Low | Pulse Low | ARTd High | NBPs Low | ARTs High | Pulse High | NBPm High | ICPm High | ARTs Low | ABPs High | awRR High | ARTm Low | NBPd High | NBPm Low | imCO2 High | awRR Low
Series 1 | 146331 | 70438 | 18761 | 7290 | 4161 | 2204 | 1108 | 1027 | 720 | 686 | 606 | 587 | 517 | 511 | 419 | 316 | 301 | 279 | 243 | 218 | 202 | 154 | 140 | 107
SpO2 Low accounts for 57% of the total of all the medium priority parameter alarms
DONE

## Slide 31
Chart: Alarm Violation Value
| 86 | 85 | 84 | 83 | 82 | 81 | 80 | 79 | 78 | 77 | 76 | 75 | 74 | 73 | 72 | 71 | 70 | 69 | 68 | 67 | 66 | 65 | 64 | 63 | 62 | 61 | 60-55 | 54-50 | 49-0
Series 1 | 2285 | 4851 | 5340 | 5513 | 5249 | 4545 | 3687 | 3087 | 2621 | 2202 | 1834 | 1548 | 1337 | 1075 | 916 | 743 | 564 | 510 | 366 | 343 | 257 | 217 | 144 | 159 | 116 | 89 | 257 | 101 | 83
Medium Priority SpO2 Parameter Alarms
Literature recommends a default value of 88 for low limit
Total SpO2 violations that occurred in NNICU with a limit set to 87 - Over a 30 days period – 64,480
Red Text

## Slide 32
Chart: Alarm Count
| CH NNICU | CH SC5 EC | CH PICU | CH SC8 | CH 8D | CH Peds ED | CH 7C | CH Peds Same D | CH PPR
Alarm Count | 141893 | 39125 | 16163 | 15430 | 14801 | 14768 | 13271 | 1283 | 1000
Unique Alarms | 13 | 9 | 38 | 6 | 14 | 17 | 12 | 15 | 10
Medium Priority Parameter Alarm Totals By Unit
Are all parameter alarms in the PICU actionable?
Data Analytics
Total Number of Unique Medium Priority Parameter Alarms By Unit
Over 30 Days – 257,734
Total Unique Parameters Alarms Types by Unit
Total Parameter Alarms by Unit
Red Text

## Slide 33
Chart: 
| HR Low | HR High | Pair PVCs | PVC Rate | Non-Sustain VT | Irregular HR | End Irregular HR | Missed Beat | Multiform PVCs | Pause | Run PVCs High | Pacer Not Pacing | Pacer Not Capture | R-on-T PVCs | Afib | End Afib | Vent Bigeminy | Vent Rhythm | etCO2 Low | SVT | Vent Trigeminy
Series 1 | 34290 | 16273 | 4084 | 2029 | 1900 | 1814 | 1394 | 1260 | 1255 | 712 | 480 | 275 | 187 | 100 | 88 | 66 | 36 | 12 | 8 | 7 | 7
Medium Priority Arrhythmia Alarms by Frequency
HR High and HR Low accounts for 76% of these alarms: can default limits be safely adjusted?
Data Analytics
Total Number of Medium Priority Arrhythmia Alarms that
Occurred Across All Units – Over 30 Days – 66,277
HR High/Low accounts for 76% of the total of the medium priority arrhythmia alarms
Red Text

## Slide 34
Data Analytics
High Priority/Red Arrhythmia Alarms
High Priority/Red Parameter Alarms
Total Number of High Priority Monitoring Alarms Across Units – Over 30 Days – 68,831
High Priority Monitoring Alarms by Frequency
Desat accounts for 73% of the total high priority alarms.
Is an Extreme Brady limit of 50 actionable?
Chart: 
| Desat | Extreme Brady | Extreme Tachy | Asystole | Apnea | VTach | Vent Fib/Tach | Brady Pulse | Tachy Pulse | ABP Disconnect | ART Disconnect | ICPm High
Series 1 | 50365 | 5545 | 3651 | 2512 | 2458 | 2168 | 1628 | 214 | 133 | 123 | 28 | 6
Red Text

## Slide 35
Data Analytics
Comparison of Leads Off to High Priority Alarms
High Priority Alarms
Variation in severity of alarm priority for Leads Off
ECRI recommends High Priority
High Priority/Red Arrhythmia Alarms
High Priority/Red Parameter Alarms
Inop Alarm for ECG Leads off – variation in priority
Chart: 
| ECG Leads Off | Desat | Extreme Brady | Extreme Tachy | Asystole | Apnea | VTach | Vent Fib/Tach | Brady Pulse | Tachy Pulse | ABP Disconnect | ART Disconnect | ICPm High
Series 1 | 23871 | 50365 | 5545 | 3651 | 2512 | 2458 | 2168 | 1628 | 214 | 133 | 123 | 28 | 6
DONE

## Slide 36
Data Analytics
Select Inop Monitoring Alarms
Leads Off is the most common and there is variation in Leads Off priority and priority for a single lead off
Chart: 
| ECG Leads Off - Low | ECG Leads Off - Medium | LL Leads Off - Medium | RA Leads Off - Medium | Check Patient ID - Medium | NBP Cuff Overpress - Low | Replace Tele Battery - Low | LA Leads Off - Medium | Insert Battery - Medium | Battery Empty - Low | Check Equipment - Medium | Transmitter Off - Low
Series 1 | 21445 | 2426 | 252 | 237 | 99 | 44 | 33 | 25 | 16 | 9 | 3 | 1
Total Number of InOp Monitoring Alarms That
Occurred Across All Units – Over 30 Days – 24,590
DONE

## Slide 37
Data Analytics
Inop/Technical Alarms by Unit
Chart: 
| CH NNICU | CH 8D | CH SC5 EC | CH 7C | CH Peds ED | CH PICU | CH SC8 | CH PPR | CH Peds Same D
Series 1 | 5117 | 4530 | 3855 | 3809 | 2616 | 2486 | 1260 | 725 | 192
Total: 24,590
DONE
Notes: Leads off is the most common for the ED and Units 30 & 40

## Slide 38
Data Analytics
AAMI task force evaluating if 20% of patients cause the majority of the alarms (65%-80%)
Patient Outliers
Can impact total alarms on a unit
Customization may have impact here
Chart: 
| 8EA | 8EA1 | 8EB | 8ED1 | 8EE1 | 8EF | 8EG | 8EG1 | 8EH | 8EH1 | 8EI1 | 8EJ | 8EJ1 | 8EK | 8EK1 | 8EL | 8EL1 | 8EM | 8EM1 | 8EN | 8EO | 8EO1 | 8EP | 8EP1 | 8EQ | 8ER | 8ES | 8ES1 | STAB1
Series 1 | 22 | 239 | 275 | 105 | 183 | 259 | 477 | 118 | 193 | 24 | 209 | 71 | 38 | 762 | 474 | 468 | 148 | 228 | 77 | 38 | 30 | 70 | 10 | 583 | 51 | 175 | 212 | 60 | 83
Room Number
NNICU, 7-Mar-2018
DONE

## Slide 39
Data Analytics
Vent Alarms by Unit
Chart: 
| MSICU | STICU
Yellow Alarms | 601 | 271
Red Alarms | 471 | 136
Yellow Vent Alarms
Red Vent Alarms
Notes: Leads off is the most common for the ED and Units 30 & 40